GOLDEN MAYSUN INT’L CO., LTD.

3Fl, No.3, Lane 105, Sec.4, Ren‐Ai Road, Taipei 106, Taiwan

Tel: 886‐2‐2731‐9118
Fax: 886‐2‐27812793



Date: July 29, 2020

Product Description

TECHNICAL SPECIFICATIONS FOR NITRILE EXAMINATION GLOVES

TECHNICAL DATA SHEET

NITRILE GLOVES‐EN 455

NITRILE GLOVE SPECIFICATION ~12 Inch, Textured, Powder Free MEDICAL EXAMINATION GLOVE

1. Length: Length of all gloves shall be not less than 300 mm for Medium and Large, and not less than 290 mm for Small. 12 Inch Gloves.

Color: Blue and White

	. Size: M 40%, L 40%, XL 20%
	
	
	
	

	
	Product
	
	Width
	
	Length

	
	Sizes
	
	(Averages)
	
	(Averages)

	
	
	
	
	
	300 mm

	2. Width& Length
	SMALL
	
	85 +/‐ 10 mm
	
	minimum

	
	
	
	
	
	300 mm

	12 Inch Glove
	MEDIUM
	
	95 +/‐ 10 mm
	
	minimum

	
	
	
	
	
	300 mm

	
	LARGE
	
	106+/‐ 10 mm
	
	minimum

	
	
	
	
	
	300 mm

	
	X‐LARGE
	
	116 +/‐ 10 mm
	
	minimum


3. Thickness: Finger‐ 0.21 mm range MIL range = 4.7 ‐ 6.7 ASTM .08mm Min. Palm ‐ 0.18 mm range MIL range = 4.3 ‐ 5.9 ASTM .08mm Min.

Cuff ‐0.15 mm range MIL range = 3.5 ‐ 4.7 ASTM .08mm Min.

PHYSICAL

4.  Waterleak: ASTM D5151, current edition shall apply as well as FDA Waterleak

Test. at 1.5 AQL so that Waterleak Testing Shall be No Greater than 2.5 upon

receipt of goods by customer unless otherwise specified by both parties.

5. Physical: Before (Average) Aging: Elongation shall be not less than 700 % Tensile Strength (MPa) shall be not less than 14.0 MPa
6. Physical: After Aging: Elongation shall be not less than 500 % (Average Tensile Strength (MPa) shall be not less than 14.0 MPa (Aging is to be at 70 +/‐2 degrees C for 166 +/‐ 2 hours.)
7. Protein Test as per ASTM D5712, current edition Not Required on Nitrile.

8. Powder Residue as per ASTM 6124, current edition. Less than 1.0 milligrams Average. (Product exceeds the current requirements.)
9. Low in Chemical Residuals supported by Electrochemistry Data under IEST‐RP‐ C005‐2 for clean room requirements. Reduced chemical irritation is goal.

10. glove is easily donned and is soft for comfortable wear. Low modulus technology. It is odor free and non‐tacky.

11. The 12 Inch Length affords the user basic protection in the medical use area as a Patient Examination Glove.

12. The NITRILE glove, is one of the best materials to replace Latex, avoiding the potential for Latex Sensitive individuals who are known to have allergic reactions and to avoid sensitizing individuals through prolonged use of latex. The Nitrile material is 100% Nitrile, and 100% Latex‐Free. It is a POWDER‐FREE Glove.

PACKAGING: 100 GLOVES (50 pair) PER DISPENSER BOX

[image: image1.jpg]\K}\ovo‘”

Giing tay y t& nhin tao khong bt Nitrile
Khong bot. Thudn ch hai tay. Chwta tiét tring
Sit dyng mt in, Nhim ngon tay

100 chie





